## 공통규격서
구매 (영문) 규격서
Commodity Description
H.S No |  | Item No. |  | Description | Quantity | Unit
|  |  |  | Mobile C-arm X-ray system | 1 | System
A. FEATURES
| 1 | . |  | Workstation complete with an articulating arm UHD monitor with touchscreen for
comfortable viewing with bigger displayed images.
| 2 | . |  | Premium C-arm with CMOS technology with ultra-eefficient crystalline structure.
| 3 | . |  | True continuous fluoroscopy captures more detailed anatomy without the distraction
of stutter, visible lag, or ghosting characteristic of “continuous pulse” fluoroscopy images
| 4 | . |  | Touch control panel with intuitive user interface for surgical team.
B. CONFIGURATION
| 1 | . |  | Generator | 1 Set
| 2 | . |  | X-ray Tube | 1 Set
| 3 | . |  | Image Receptor | 1 Set
| 4 | . |  | Collimator | 1 Set
| 5 | . |  | Touch control panel | 1 Set
| 6 | . |  | Image Display | 1 Set
| 7 |  |  | Image Processing
| 8 | . |  | Image Achieving system | 1 Set
| 9 | . |  | Vascular imaging | 1 Set
| 10 | . |  | Vascular MTS | 1 Set
| 11 | . |  | Physical movements | 1 Set
| 12 | . |  | System Accessories | 1 Set
| 13 | . |  | Vender Unique options
C. SPECIFICATIONS
| 1 | . |  | Generator | 1 Set
1) Type : 18 kHz high frequency or equivalent
2) Power : 15 kW power or more
3) KV range : Up to 120 kvp
4) Continuous high level fluoro or equivalent
5) Pulsed high level fluoro or equivalent
6) Digital spot or equivalent
7) Digitial Cine Pulse Mode or equivalent
8) High level fluoro or High Level Mode or equivalent
9) Split-block generator or mono-block with Energy Storage Unit
10) Uninterruptible Power Supply or equivalent
| 2 | . |  | X-ray Tube | 1 Set
|  |  | 1) | Type : Rotating Anode
|  |  | 2) | Focal spots : 0.3 and 0.6 mm or 0.3 and 0.5 mm
|  |  | 3) | Anode heat capacity : 300,000 HU or more
|  |  | 4) | Anode Cooling rate : 85,000 HU/min or more
|  |  | 5) | Housing cooling rate : 268W or more
| 3 | . |  | Flat Panel Detector | 1 Set
|  |  | 1) | Type : CMOS Flat Panel Detector (CFD)
|  |  | 2) | Field of View : 31 cm / 21 cm / 15 cm or equivalent
|  |  | 3) | Active Matirx : 1548 x 1524 pixels or more
|  |  | 4) | Norminal resolution : 2.9 lp/mm or more
|  |  | 5) | DQE(0) : 70% or more
|  |  | 6) | Pixel Pitch : 152 μm or more
|  |  | 7) | Anti-scatter Removable grid or equivalent
| 4 | . |  | Collimator | 1 Set
|  |  | 1) | On-screen collimator position indication or equivalent
|  |  | 2) | PreView iris collimator or equivalent
|  |  | 3) | PreView Tungsten rotation double leaf collimator or equivalent
|  |  | 4) | Adjust collimators without X-ray exposure or equivalent
| 5 | . |  | Touchscreen control panel | 1 Set
|  |  | 1) | Touchscreen panel on C-arm stand or equivalent
|  |  | 2) | Resolution : 640 x 480 or equivalent
|  |  | 3) | Intuitive user interface for surgeon or equivalent
|  |  | 4) | Tilt and Swivel function or equivalent
| 6 | . |  | Image Display | 1 Set
|  |  | 1) | 32" Ultra high definition (UHD) Color Display or 19" TFT monitors
|  |  | 2) | Display mounted on an articulating arm on workstation or ceiling mount
|  |  | 3) | Resolution : 3840 x2160 or 1280x1024(2ea)
|  |  | 4) | Maximum brightness : 600 cd/m2 or more
|  |  | 5) | Touch Screen system control or equivalent
| 7 | . |  | Image Processing | 1 Set
|  |  | 1) | Image Processing :
|  |  | 2) | Preset Imaging Profiles
|  |  | 3) | Noise filter with on-screen indicator
|  |  | 4) | Minimal difference spatiotemporal noise filter (MDST) or equivalent
|  |  | 5) | Real-time dynamic general purpose range management (GDRM) or equivalent
|  |  | 6) | Automatic digital brightness and contrast control or equivalent
|  |  | 7) | Manual digital brightness and contrast control or equivalent
|  |  | 8) | Negate mode or equivalent
|  |  | 9) | Save and auto-save feature or equivalent
|  |  | 10) | Swap and auto-swap feature or equivalent
|  |  | 11) | Last image hold or equivalent
|  |  | 12) | Image Storage : 40,000 or more
|  |  | 13) | USB export with DICOM viewer for displaying images on PC platforms
|  |  | 14) | Livezoom for 4 times magnified FOV without dose increase and enhanced Noise Reduction(eNR)
or equivalent
|  |  | 15) | On-screen free hand image marking or Live Graphical Overlay or Anatomiacal Marking Tool
|  |  | 16) | HIPAA SecureView or equivalent
|  |  | 17) | Zoom and pan function
|  |  | 18) | Image annotation
|  |  | 19) | Measurement
| 8 | . |  | Image Achieving system | 1 Set
|  |  | 1) | Integrated DICOM interface or equivalent
- Storage commit, store, query, retrieve, modality, worklist and print
- Queue DICOM Store, Storage commit and Print requests for later transmission
2 )Additional achieve system or equivalent
- USB - three ports
| 9 | . |  | Vascular imaging | 1 Set
|  |  | 1) | Vascular preset imaging profile or equivalent
|  |  | 2) | Real-time digital subtraction (DSA)
|  |  | 3) | Peak opacification or equivalent
|  |  | 4) | Roadmapping
|  |  | 5) | Re-registration or equivalent
|  |  | 6) | Variable landmarking or equivalent
|  |  | 7) | Mask save/recall
|  |  | 8) | Noise Reduction Filter
| 10 | . |  | Vascular MTS | 1 Set
|  |  | 1) | Bolus chase preset imaging profile with Motion Tolerant Subtraction (MTS)or equivalent
|  |  | 2) | Digital Cine pulse mode or equivalent
- 15, 30 pulses/sec or equivalent
- Up to 150mA or equivalent
- 9 ms pulse width or equivalent
|  |  | 3) | Up to 30 fps Cine or equivalent
- Recording/playback rate : 4, 8, 15, 30 fps or equivalent
- Minimum recording time : 60 minutes or equivalent
- Automatic image playback
- Frame-by-Frame review
| 11 | . |  | Physical movements | 1 Set
|  |  | 1) | SID range : 99 cm or more
|  |  | 2) | Free space in arc : 77 cm or more
|  |  | 3) | Depth in arc : 67 cm or more
|  |  | 4) | Orbital rotation : 145 deg or more
|  |  | 5) | Lateral rotation : 360 deg or more
|  |  | 6) | Additional L shape articulation or equivalemt
|  |  | 7) | Wig/wag : 12 deg or more
|  |  | 8) | Horizontal travel : 20 cm
|  |  | 9) | Vertical travel : 45 cm or more
|  |  | 10) | Multi-functional controls or equivalent
(a) Footswitch
(b) Hand-held control
|  |  | 11) | Long Range Hand Rail or Additional Hand Rail for Low Lateral Positiong or equivalent
| 12 | . |  | System Accessories | 1 Set
|  |  | 1) | Operator Manual
|  |  | 2) | Integrated Laser Aimer
|  |  | 3) | Carbon table
| 13 | . |  | Vender Unique options | 1 Set
A Type
|  |  | 1) | Articulated monitor display for optimal viewing angle without workstation
repositioning
|  |  | 2) | Tube-Generator split block for larger FOV and collision protection via smaller profile
|  |  | 3) | Digital zoom : Up to 4 times of digital zoom without enhancing mA and kV
|  |  | 4) | Continuous fluoroscopy for motion artifact eliminated imaging
|  |  | 5) | Additional L shaped C-arm joint for extended angulation and oblique imaging
|  |  | 6) | eNR (Enhanced Noise Reduction) Filter
B Type
|  |  | 1) | Power : 25kW
|  |  | 2) | High Level Mode
|  |  | 3) | Laser Crosshair
|  |  | 4) | Dual 19" monochrome TFT
|  |  | 5) | Stenosis quantification : Quantification program for geometric and densitometric values
|  |  | 6) | Energy Storage Unit for up to 3000W fluoro high level mode or equivalent
|  |  | 7) | Motorized Monitor height adjust
|  |  | 8) | Live Graphical Overlay with Draw AAA
|  |  | 9) | UPS 5 min of monitor cart
Remarks
1. Installation
Installation of this equipment at customer's premise should be done on the contract's responsibility.
2. Warranty Conditions
Contractor or Supplier should be responsible free of charge after service(A/S) whenever
there is any defect or trouble during 3 years after the equipment is officially accepted by end-user.
3. In no case shall the amount of damages to be paid by Seller to the Purchaser exceed the
contract price.